## Overview
Overview
Connection
USB port defition (Virtual COM port) |  | Uart port definition
Baudrate | 115200 | Baudrate | 38400
Parity | None | Parity | None
Stop Bits | One | Stop Bits | One
Data bits | 8 | Data bits | 8
Command Flow
See sheet "Command Flow" for complete overview of communication commands
Communication Protocol
See sheet "Communication Protocol" for complete overview of communication commands
Definitions
CRC - Cyclic Redundancy Check
Commands are appended with a CRC sum to detect any accidental changes to raw data.
Use the following standard to calculate CRC sums (code algorithm can be found on "CRC" sheet):
Name | CRC-16-IBM
Seed | 0xA001
Initial Value | 0x0000
Status Byte Definitions
Status Byte #1 |  | Status Byte #3
Bit 0 (LSB) equal 1: Not all hardware available |  | empty
Bit 1 equal 1: No or faulty EEPROM
Bit 2 equal 1: No temperature limits received (for controlled mode)
Bit 3 equal 1: Cannot reach lens focal power (Focal Power Controlled)/position (Position Controlled)
Bit 4 equal 1: Focal power inversion (defined by user set temperature range)
Bit 5 equal 1: Temperature is outside product specifications
Bit 6 equal 1: Focal power out of guaranteed range (defined by user set temperature range)
Bit 7 (MSB) equal 1: Temp out of range specified by user
Status Byte #2 |  | Status Byte #4
Bit 0 (LSB) equal 1: The connected lens is not compatible with the firmware on the lensdriver. |  | empty
Commands that return status bytes (previously known as error byte)
Status Bytes | command
Status Byte #1 | #0501 Temperature reading
#0503 Get Status
#1100 Set Gain Variable
#0308 Change to Focal Power Controlled Mode
#0309 Get Temperature Limits
Status Byte #2 | #0503 Get Status
Status Byte #3 | #0503 Get Status
Status Byte #4 | #0503 Get Status

## Command Flow
Command Flow
Boot procedure
Start command (#0101)
Current mode
Set Current mode (#0101)
DC Mode Commands
Get/set current(#0201)
Mode independent
commands
Mode independent
commands
Set signal mode(#0301, #0302, #0304)
Signal Mode Commands
Get/set Lower swing limit (#0305)
Get/set Upper swing limit (#0306)
Get/set frequency(#0307)
Signal mode
Controlled Mode Commands
Get/set focal power(#0310)
Set temperature limits (#0309)
Set focal power controlled mode (#0308)
Focal power controlled mode
Mode independent
commands
Mode Independent Commands
Get lens temperature (#0501)
Analog input reading(#1001)
Get firmware version(#0701)
Set analog/serial control(#0601-#0602)
Get/set gain variable(#1100)
Get/set max output current (#0401)
Get/set software current upper limit (#0402)
Get/set software current lower limit (#0403)
Device ID read (#0901)
Position controlled mode
Set position controlled Mode (#0320)
Position Controlled Mode Commands
Get/set desired position(#0311)
Sensor reading (#1002)

## Communication Protocol
Communication Protocol
Commands marked as deprecated will be removed in the next release.
[9 empty] | Send Commands | [9 empty] | Reply Commands
Categories | Name | Internal | Command Status | LD4 Compatible Firmware Types | LD4 Firmware Version Type A | ID Number | Short Description | Bidirectional | Send Command: Is sent from the PC to the Device | <Prefix> (ASCII) | <Read/Write> (ASCII) | <Coding> (ASCII) | <Channel> (ASCII) | <DataIn> (Bytes) | Data in encoding For Firmware Type A | Data in encoding For Firmware Type B,C and E | <CRC16> 2 bytes: Low Byte High Byte | Example Send Command Red: ASCII Characters <Blue>: Binary Values | Reply Command: Is sent from the Device to the PC | <Prefix> (ASCII) | <Coding> (ASCII) | <Channel> (ASCII) | <DataOut> (Bytes) | Data Out decoding For Firmware Type A | Data Out decoding For Firmware Type B, C and E | <CRC16> 2 bytes: Low Byte High Byte | <end> (ASCII) | Example Reply Command Red: ASCII Characters <Blue>: Binary Values | Comments
01-Firmware Commands and Properties | Handshake | No | ACTIVE | A,B,C,D,E | 1.8 | #0101 | returns a string if the system is working | yes | <PrefixIn> | Start | none | none | none | none |  |  | none | START | <PrefixOut><end> | "Ready" | none | none | none |  |  | none | "\r\n" | READY\r\n
01-Firmware Commands and Properties | Firmware Type | No | ACTIVE | A,B,C,D,E | 1.8 | #0103 | returns the firmware type of the compiled firmware | yes | <Prefix><crc16> | "H" | none | none | none | none |  |  | LH | Example of Firmware Type command: H<CRC> Hex: 0x48 0x00 0x36 | <Prefix><DataOut><crc16><end> | "H" | none | none | <1 byte firmware type> |  |  | LH | "\r\n" | H<Firmware Type><CRC>\r\n | There are several firmware types that differ in the functionality that is available. This command makes up part of the versioning scheme <major version>.<minor version>.<build>.<revision> (<ID>/<branch>-<firmware type>).
01-Firmware Commands and Properties | Firmware Branch | No | ACTIVE | A,B,C,D,E | 1.8 | #0104 | returns the firmware branch of the compiled firmware. | yes | <Prefix><crc16> | "F" | none | none | none | none |  |  | LH | Example of Firmware Branch command: F<CRC> Hex: 0x46 0x81 0xF2 | <Prefix><DataOut><crc16><end> | "F" | none | none | <1 byte firmware branch> | 0 - Trunk 1 - Branch 1 2 - Branch 2 | 0 - Trunk 1 - Branch 1 2 - Branch 2 | LH | "\r\n" | F<Firmware Branch><CRC>\r\n | This command makes up part of the versioning scheme <major version>.<minor version>.<build>.<revision> (<ID>/<branch>-<firmware type>).
01-Firmware Commands and Properties | Part Number Command | No | ACTIVE | A,B,C,D,E | 1.8 | #0105 | returns the part number of the compiled firmware. | yes | <Prefix><crc16> | "J" | none | none | none | none |  |  | LH | Example of Part Number command: J<CRC> Hex: 0x4A 0x81 0xF7 | <Prefix><DataOut><crc16><end> | "J" | none | none | <3 bytes part number> |  |  | LH | "\r\n" | J<Part Number><CRC>\r\n | Returns the first 6 digits of the Part number. The internal part number describing the firmware. This Number identifies the project the firmware belongs to. This command makes up part of the versioning scheme <major version>.<minor version>.<build>.<revision> (<part number>/<branch>-<firmware type>).
01-Firmware Commands and Properties | Device ID - Read | No | ACTIVE | A,B,C,D,E | 1.8 | #0901 | Send 8 dummy bytes and read the ID from ID memory | yes | <Prefix><DataIn><crc16> | "IR" | none | none | none | <8 dummy bytes> |  |  | LH | Example Set ID to 00000000: IR<00000000><CRC> Hex: 0x49 0x52 0x00 0x00 0x00 0x00 0x00 0x00 0x00 0x00 0x37 0x98 | <Prefix><DataOut><crc16><end> | "IR" | none | none | <8 bytes unsigned integer Device ID> |  |  | LH | "\r\n" | IR<device ID><CRC>\r\n
01-Firmware Commands and Properties | Firmware Version Read | No | DEPRECATED | A,B,C,D,E | 1.8 | #0701 | Read out the firmware version including the Major Revison, Minor Revision, Build and Revison | yes | <Prefix><DataIn><crc16> | "V" | none | none | none | <1 dummy byte> |  |  | LH | Example Read Version Command: V<0><CRC> Hex: 0x56 0x00 0x3F 0xA0 | <Prefix><DataOut><crc16><end> | "V" | none | none | <1 byte unsigned integer Firmware Major ><1 byte unsigned integer Firmware Minor> <2 bytes unsigned integer Build><2 bytes unsigned integer Revison > |  |  | LH | "\r\n" | V<DataOut><CRC>\r\n | This command makes up part of the versioning scheme <major version>.<minor version>.<build>.<revision> (<ID>/<branch>-<firmware type>). The firmware version makes up the major version, minor version, build and revision.
01-Firmware Commands and Properties | Firmware Version Read | No | NEW | A,B,C,D,E | 1.9 | #0701 | Read out the firmware version including the Major Revison, Minor Revision, Build and Revison | yes | <Prefix><crc16> | "V" | none | none | none | none |  |  | LH | Example Read Version Command: V<CRC> Hex: 0x56 0x80 0x3E | <Prefix><DataOut><crc16><end> | "V" | none | none | <1 byte unsigned integer Firmware Major ><1 byte unsigned integer Firmware Minor> <2 bytes unsigned integer Build><2 bytes unsigned integer Revison > |  |  | LH | "\r\n" | V<DataOut><CRC>\r\n | This command makes up part of the versioning scheme <major version>.<minor version>.<build>.<revision> (<ID>/<branch>-<firmware type>). The firmware version makes up the major version, minor version, build and revision.
01-Firmware Commands and Properties | Set Upper Software Current Limit | No | ACTIVE | A,B,C,D,E | 1.8 | #0402 | Set the upper software current limit in channel A. This limit cannot be higher than the maximum output current. | yes | <Prefix><Write><Coding><Channel><DataIn><crc16> | "C" | "w" | "U" | "A" | <2bytes signed integer upper current limit> | Encoded value = upper limit * (4095 /max current), where the default max current is 293mA (command #0401). Encoded values can only be within the range -4095 to 4095 | Encoded value = upper limit * (4095 /max current), where the default max current is 293mA (command #0401). Encoded values can only be within the range -4095 to 4095 | LH | Example Set Upper Limt to 292mA: CwUA<4095><CRC> Hex: 0x43 0x77 0x55 0x41 0x0F 0xFF 0xFE 0x90 | <Prefix><Coding><Channel><DataOut><crc16><end> | "C" | "U" | "A" | <2 bytes signed integer upper current limit> | Upper limit = encoded value * (max current/4095), where the default max current is 293mA (command #0401). | Upper limit = encoded value * (max current/4095), where the default max current is 293mA (command #0401). | LH | "\r\n" | CUA<encoded upper current limit><CRC> \r\n | Changing this value will effect the range of achiveable current and focal power values.
01-Firmware Commands and Properties | Get Upper Software Current Limit | No | ACTIVE | A,B,C,D,E | 1.8 | #0402 | Get the upper software current limit in channel A. This limit cannot be higher than the maximum output current. | yes | <Prefix><Read><Coding><Channel><DataIn><crc16> | "C" | "r" | "U" | "A" | <2 dummy bytes> |  |  | LH | Example Get the Upper Limt: CrUA<0><0><CRC> Hex: 0x43 0x72 0x55 0x41 0x00 0x00 0x77 0x20 | <Prefix><Coding><Channel><DataOut><crc16><end> | "C" | "U" | "A" | <2 bytes signed integer upper current limit> | Upper limit = encoded value * (max current/4095), where the default max current is 293mA (command #0401). | Upper limit = encoded value * (max current/4095), where the default max current is 293mA (command #0401). | LH | "\r\n" | CUA<encoded upper current limit><CRC> \r\n
01-Firmware Commands and Properties | Set Lower Software Current Limit | No | ACTIVE | A,B,C,D,E | 1.8 | #0403 | Set the lower software current limit in channel A. This limit cannot be higher than the maximum output current. | yes | <Prefix><Write><Coding><Channel><DataIn><crc16> | "C" | "w" | "L" | "A" | <2 bytes signed integer min current> | Encoded value = lower limit * (4095 /max current), where the default max current is 293mA (command #0401). Encoded values can only be within the range -4095 to 4095 | Encoded value = lower limit * (4095 /max current), where the default max current is 293mA (command #0401). Encoded values can only be within the range -4095 to 4095 | LH | Example Set Lower Limt to 0mA: CwLA<0><CRC> Hex: 0x43 0x77 0x4C 0x41 0x00 0x00 0xBC 0x7C | <Prefix><Coding><Channel><DataOut><crc16><end> | "C" | "L" | "A" | <2 bytes signed integer lower current limit> | Lower limit = encoded value * (max current/4095), where the default max current is 293mA (command #0401). | Lower limit = encoded value * (max current/4095), where the default max current is 293mA (command #0401). | LH | "\r\n" | CLA<encoded lower current limit><CRC> \r\n | Changing this value will effect the range of achiveable current and focal power values.
01-Firmware Commands and Properties | Get Lower Software Current Limit | No | ACTIVE | A,B,C,D,E | 1.8 | #0403 | Get the lower software current limit in channel A. This limit cannot be higher than the maximum output current. | yes | <Prefix><Read><Coding><Channel><DataIn><crc16> | "C" | "r" | "L" | "A" | <2 dummy bytes> |  |  | LH | Example Get the Lower Limt: CrLA<0><0><CRC> Hex: 0x43 0x72 0x4C 0x41 0x00 0x00 0x70 0x7C | <Prefix><Coding><Channel><DataOut><crc16><end> | "C" | "L" | "A" | <2 bytes signed integer lower current limit> | Lower limit = encoded value * (max current/4095), where the default max current is 293mA (command #0401). | Lower limit = encoded value * (max current/4095), where the default max current is 293mA (command #0401). | LH | "\r\n" | CLA<encoded lower current limit><CRC> \r\n
01-Firmware Commands and Properties | Set Gain Variable | No | ACTIVE | A | 1.8 | #1100 | Set the gain variable for focal power drift compensation. The set command responds with the achievable focal power range at the given temperature limits and given gain variable. The get command simply returns the gainvariable. Gain variable must be between 0 and 5. | yes | <Prefix><Write><DataIn><crc16> | "O" | "w" | none | none | <2 bytes unsigned integer Gain Variable> | Encoded value = gain variable * 100 | Encoded value = gain variable * 100 | LH | Example Set Gain to 1 Command: Ow<100> <CRC> Hex: 0x4F 0x77 0x00 0x64 0xA7 0x25 | <Prefix><Coding><DataOut><crc16><end> | "O" | "w" | none | <1 byte unsigned integer ErrorByte> <2 bytes signed integer max focal power limit><2 bytes signed integer min focal power limit> | Focal power = (encoded value/200) -5 | Focal power = encoded value/200 | LH | "\r\n" | Ow<errorbyte><encoded max focal power limit><encoded min focal power limit><CRC> \r\n | Error byte defniition see sheet "Help"
01-Firmware Commands and Properties | Get Gain Variable | No | ACTIVE | A | 1.8 | #1100 | Get the gain variable for focal power drift compensation. The set command responds with the achievable focal power range at the given temperature limits and given gain variable. The get command simply returns the gainvariable. Gain variable must be between 0 and 5. | yes | <Prefix><Read><DataIn><crc16> | "O" | "r" | none | none | <2 dummy bytes> |  |  | LH | Example Get Gain Variable: Or<dummy><CRC> Hex: 0x4F 0x72 0x00 0x00 0xB6 0xCF | <Prefix><Coding><DataOut><crc16><end> | "O" | "r" | none | <2 bytes unsigned integer Gain Variable> | Gain variable = encoded value / 100 | Gain variable = encoded value / 100 | LH | "\r\n" | Or<Gain variable><CRC> \r\n
02-Error Commands | Communication Error | No | ACTIVE | A,B,C,D,E | 1.8 | #0100 | Return an error message to the user. | No | none | none | none | none | none | none |  |  | none |  | <Prefix><DataOut><end> | "E" | none | none | <1 byte unsigned integer Command Error ID> | ID 1 - CRC failed 2 - Command not available in firmware type 3 - Command not recognized 4 - Lens is not compatible with firmware | ID 1 - CRC failed 2 - Command not available in firmware type 3 - Command not recognized 4 - Lens is not compatible with firmware | LH | "\r\n" | E<Command Error><CRC>\r\n ex.: send: St4rt receive: E<error id><crc>\r\n | If the command ID write (#0901) is sent in a non-internal version of the firmware. Then the Communication Error command will be returned. This command will include the Command Error of 001 indicating this is an internal command, along with the received command (i.e. I W 0000 0000 CRC for this example). If a command is sent and there is a CRC error with the sent command a nack (N/r/n) will be returned to the user.
03-Lens Commands and Properties | Serial Number | No | ACTIVE | A,B,C,D,E | 1.8 | #0102 | returns the serial number of the connected lens | yes | <Prefix><crc16> | "X" | nonne | none | none | none |  |  | LH | Example of Serial Number command: X <CRC> Hex: 0x58 0x01 0xFA | <Prefix><DataOut><crc16><end> | "X" | none | none | <8 bytes serial number> |  |  | LH | "\r\n" | X<Serial Number><CRC>\r\n | If no EEPROM has been read in then the command returnsthe reply command with eight 0's for the value of the serial number.
03-Lens Commands and Properties | Set Current | No | ACTIVE | A,B,C,D,E | 1.8 | #0201 | Get/set current to channel A | no | <Channel><Write><DataIn><crc16> | "A" | "w" | none | none | <2 bytes signed integer current> | Encoded value = current * (4095 /max current), where the default max current is 293mA (command #0401). Encoded current values can only be within the range -4095 to 4095 | Encoded value = current * (4095 /max current), where the default max current is 293mA (command #0401). Encoded current values can only be within the range -4095 to 4095 | LH | Example Current Set of 100 mA: Aw<1402><CRC> Hex: 0x41 0x77 0x05 0x7A 0x26 0x95 | none | none | none | none | none | none | none | none | none | none
03-Lens Commands and Properties | Get Current | No | ACTIVE | A,B,C,D,E | 1.8 | #0201 | Get/set current to channel A | no | <Channel><Read><DataIn><crc16> | "A" | "r" | none | none | <2 dummy bytes> |  |  | LH | Example Read Current Command: Ar<2 dummy bytes><CRC> Hex: 0x41 0x72 0x00 0x00 0xb4 0x27 | <Channel><DataOut><crc16><end> | none | none | "A" | <2 bytes signed integer current> | Current = encoded value * (max current/4095), where the default max current is 293mA (command #0401). | Current = encoded value * (max current/4095), where the default max current is 293mA (command #0401). | LH | "\r\n" | A<encoded current><CRC>\r\n | Commands to set current are sent as one char, followed by two bytes containing the 16bit *signed*current value (can only be -4095 to 4095) followed by two bytes for the CRC16 sum. If a command is out of boundaries (i.e. 5000) it is limited to 4095 by the firmware, no overflow will occur. If the software current limit is set lower than 4095 (i.e. 3000, see calibration commands), current requests bigger than the software limit will be reduced down to the limit (i.e. 3000).
03-Lens Commands and Properties | Get Status | No | NEW | A,B,C,D,E | 1.9 | #0503 | Return the firmware status information. This commands was the status byte on the Temperature Reading command. | yes | <Prefix><Channel><crc16> | "S" | "r" | none | none | none |  |  | LH | Example Get Status Command: Sr<CRC> Hex: 0x53 0x72 0xBC 0xD5 | <Prefix><DataOut><crc16><end> | "S" | none | none | <status byte #4><status byte #3><status byte #2><status byte #1> | Status byte definition can be found on overview sheet | Status byte definition can be found on overview sheet | LH | "\r\n" | S<4 status bytes><CRC> \r\n | For error byte defniition see sheet "Overview".
03-Lens Commands and Properties | Temperature Reading | No | DEPRECATED | A,B,C,D,E | 1.8 | #0501 | Returns the temperature reading of the lens of channel A DEPRECATED. This command will be remove in the next release. | yes | <Prefix><Channel><crc16> | "T" | none | none | "A" | none |  |  | LH | Example Temperature Read Command: TA<CRC> Hex: 0x54 0x41 0xFE 0xF0 | <Prefix><Channel><Coding><DataOut><crc16><end> | "T" | none | "A" | <1 byte unsigned integer status byte> <2 bytes signed integer temperature> | Temperature in °C = <16 bit signed integer temperature> * 0.0625 Status byte definition can be found on overview sheet | Temperature in °C = <16 bit signed integer temperature> * 0.0625 | LH | "\r\n" | TA<status byte> <encoded temperature><CRC> \r\n | The temperature must be between 120 and -120 degrees celsius. Error byte defniition see sheet "Help". If the temperature sensor is not connected then 4095 will be returned to the user. Once a sensor is connected it can be read, even if not all the hardware is available for that firmware type.
03-Lens Commands and Properties | Temperature Reading | No | NEW | A,B,C,D,E | 1.9 | #0501 | Returns the temperature reading of the lens of channel A | yes | <Prefix><Channel><crc16> | "TC" | none | none | "A" | none |  |  | LH | Example Temperature Read Command: TCA<CRC> Hex: 0x54 0x43 0x41 0xB0 0xD0 | <Prefix><Channel><DataOut><crc16><end> | "TC" | none | "A" | <2 bytes signed integer temperature> | Temperature in °C = <16 bit signed integer temperature> * 0.0625 | Temperature in °C = <16 bit signed integer temperature> * 0.0625 | LH | "\r\n" | TCA<encoded temperature><CRC> \r\n | The temperature must be between 120 and -120 degrees celsius. If the temperature sensor is not connected then 4095 will be returned to the user. Once a sensor is connected it can be read, even if not all the hardware is available for that firmware type.
03-Lens Commands and Properties | Analog Input Reading | No | ACTIVE | A,B,C,E | 1.8 | #1001 | Get analog reading of Channel A | yes | <Prefix><Channel><crc16> | "GA" | none | none | "A" | none |  |  | LH | Example Analog Input Command: GAA<CRC> Hex: 0x47 0x41 0x41 0x40 0x75 | <Prefix><Channel><DataOut><crc16><end> | "GA" | none | "A" | <2 bytes unsigned integer Analog input value> |  |  | LH | "\r\n" | GAA<analog input value><CRC>\r\n
03-Lens Commands and Properties | Set Signal Generator Upper Current Limit | No | ACTIVE | A,B,C,D,E | 1.8 | #0305 | Set the upper current swing limit for channel A. Command can only be used in sinusoidal, rectangular, and triangular mode. | no | <Prefix><Write><Coding><Channel><DataIn><crc16> | "P" | "w" | "U" | "A" | <2 bytes signed integer upper swing limit><2 dummy bytes> | Encoded value = upper swing limit * (4095 /max current), where the default max current is 293mA (command #0401). Encoded values can only be within the range -4095 to 4095 | Encoded value = upper swing limit * (4095 /max current), where the default max current is 293mA (command #0401). Encoded values can only be within the range -4095 to 4095 | LH | Example Upper Swing Set to 100mA: PwUA<1402><0><0><CRC> Hex: 0x50 0x77 0x55 0x41 0x05 0x7A 0x00 0x00 0x13 0x24 | none | none | none | none | none | none | none | none | none | none | The upper swing limit for current is used for Trinangular, Square and Sinusoidal mode.
03-Lens Commands and Properties | Get Signal Generator Upper Current Limit | No | ACTIVE | A,B,C,D,E | 1.8 | #0305 | Get the upper current swing limit for channel A. Command can only be used in sinusoidal, rectangular, and triangular mode. | yes | <Prefix><Read><Coding><Channel><DataIn><crc16> | "P" | "r" | "U" | "A" | <2 dummy bytes> |  |  | LH | Example Read Upper Swing Limit: PrUA<1402><0><0><CRC> Hex: 0x50 0x72 0x55 0x41 0x00 0x00 0x75 0x83 | <Prefix><channel><Coding><DataOut><crc><end> (where perfix is 2 bytes 'MP') | "P" | "U" | "A" | <2 bytes signed integer upper swing limit> | Upper swing limit = encoded value * (max current/4095), where the default max current is 293mA (command #0401). | Upper swing limit = encoded value * (max current/4095), where the default max current is 293mA (command #0401). | LH | "\r\n" | PUA<encoded upper swing limit><0><0><CRC>\r\n | The upper swing limit for current is used for Trinangular, Square and Sinusoidal mode.
03-Lens Commands and Properties | Set Signal Generator Lower Current Limit | No | ACTIVE | A,B,C,D,E | 1.8 | #0306 | Set the lower current swing limit for channel A. Command can only be used in sinusoidal, rectangular, and triangular mode. | no | <Prefix><Write><Coding><Channel><DataIn><crc16> | "P" | "w" | "L" | "A" | <2 bytes signed integer lower swing limit><2 dummy bytes> | Encoded value = upper swing limit * (4095 /max current), where the default max current is 293mA (command #0401). Encoded values can only be within the range -4095 to 4095 | Encoded value = upper swing limit * (4095 /max current), where the default max current is 293mA (command #0401). Encoded values can only be within the range -4095 to 4095 | LH | Example Lower Swing Set to 50mA: PwLA<701><0><0><CRC> Hex: 0x50 0x77 0x4C 0x41 0x02 0xBD 0x00 0x00 0xA1 0xA4 | none | none | none | none | none | none | none | none | none | none | The lower swing limit for current is used for Trinangular, Square and Sinusoidal mode.
03-Lens Commands and Properties | Get Signal Generator Lower Current Limit | No | ACTIVE | A,B,C,D,E | 1.8 | #0306 | Get the lower current swing limit for channel A. Command can only be used in sinusoidal, rectangular, and triangular mode. | yes | <Prefix><Read><Coding><Channel><DataIn><crc16> | "P" | "r" | "L" | "A" | <2 dummy bytes> |  |  | LH | Example Read Lower Swing Limit: PrLA<701><0><0><CRC> Hex: 0x50 0x72 0x4C 0x41 0x00 0x00 0x72 0xDF | <Prefix><channel><Coding><DataOut><crc><end> (where perfix is 2 bytes 'MP') | "P" | "L" | "A" | <2 bytes signed integer lower swing limit> | Lower swing limit = encoded value * (max current/4095), where the default max current is 293mA (command #0401). | Lower swing limit = encoded value * (max current/4095), where the default max current is 293mA (command #0401). | LH | "\r\n" | PLA<encoded lower swing limit><0><0><CRC>\r\n | The lower swing limit for current is used for Trinangular, Square and Sinusoidal mode.
03-Lens Commands and Properties | Set Signal Generator Frequency | No | ACTIVE | A,B,C,D,E | 1.8 | #0307 | Set the channel A frequency. Command can only be used in sinusoidal, rectangular, and triangular mode. | no | <Prefix><Write><Coding><Channel><DataIn><crc16> | "P" | "w" | "F" | "A" | <4 bytes unsigned integer frequency> | Encoded value = frequency * 1000 | Encoded value = frequency * 1000 | LH | Example Frequency Set to 100 Hz: PwFA<100000><CRC> Hex: 0x50 0x77 0x46 0x41 0x00 0x01 0x86 0xA0 0x03 0x4A | none | none | none | none | none | none | none | none | none | none | The frequency is used for Trinangular, Square and Sinusoidal mode. 0.1hz ( 100 as the coded value) is the minimum frequency.
03-Lens Commands and Properties | Get Signal Generator Frequency | No | ACTIVE | A,B,C,D,E | 1.8 | #0307 | Get the channel A frequency. Command can only be used in sinusoidal, rectangular, and triangular mode. | yes | <Prefix><Read><Coding><Channel><DataIn><crc16> | "P" | "r" | "F" | "A" | <4 dummy bytes> |  |  | LH | Example Get Frequency: PwFA<0><0><0><0><CRC> Hex: 0x50 0x72 0x46 0x41 0x00 0x00 0x00 0x00 0x65 0x92 | <Prefix><channel><Coding><DataOut><crc><end> (where perfix is 2 bytes 'MP') | "P" | "F" | "A" | <4 bytes unsigned integer frequency> | Frequency = encoded value / 1000 | Frequency = encoded value / 1000 | LH | "\r\n" | PFA<encoded frequency><0><0><CRC>\r\n | The frequency is used for Trinangular, Square and Sinusoidal mode.
03-Lens Commands and Properties | Set Temperature Limits | No | ACTIVE | A,B,C,D,E | 1.8 | #0309 | Set the upper and lower temperature limits to channel A. The command responds with the achievable focal power range at the given temperature limits. | no | <Prefix><Write><Coding><Channel><DataIn><crc16> | "P" | "w" | "T" | "A" | <2 bytes signed integer maximum temperature><2 bytes signed integer minimum temperature> | Encoded value = temperature * 16 | Encoded value = temperature * 16 | LH | Example focal power Limit Set to 20-45 °C Range: PwTA<720><320><CRC> Hex: 0x50 0x77 0x54 0x41 0x02 0xD0 0x01 0x40 0x33 0xC1 | <Prefix><Coding><DataOut><crc16><end> | "P" | "T" | none | <1 byte unsigned integer ErrorByte><2 bytes signed integer max focal power limit><2 bytes signed integer min focal power limit> | Focal power = (encoded value/200) -5 | Focal power = encoded value/200) | LH | "\r\n" | PT<errorbyte><encoded max focal power limit><encoded min focal power limit><CRC> \r\n | The temperature must be between 120 and -120 degrees celsius. Lens need to be connected. Error byte defniition see sheet "Help"
03-Lens Commands and Properties | Get Temperature Limits | No | ACTIVE | A,B,C,D,E | 1.8 | #0309 | Get the upper and lower temperature limits to channel A. The command responds with the achievable focal power range at the given temperature limits. | yes | <Prefix><Read><Coding><Channel><DataIn><crc16> | "P" | "r" | "T" | "A" | <4 dummy bytes> |  |  | LH | Example focal power Limit Set to 20-45 °C Range: PrTA<0><0><0><0><CRC> Hex: 0x50 0x72 0x54 0x41 0x00 0x00 0x00 0x00 0x66 0xE0 | <Prefix><Coding><DataOut><crc16><end> | "P" | "T" | none | <1 byte unsigned integer ErrorByte><2 bytes unsigned integer max focal power limit><2 bytes unsigned integer min focal power limit> | Focal power = (encoded value/200) -5 | Focal power = encoded value/200) | LH | "\r\n" | PT<errorbyte><encoded max focal power limit><encoded min focal power limit><CRC> \r\n
03-Lens Commands and Properties | Set Max Output Current | No | ACTIVE | A,B,C,D,E | 1.8 | #0401 | Set the maximum firmware output current in channel A. | yes | <Prefix><Write><Coding><Channel><DataIn><crc16> | "C" | "w" | "M" | "A" | <2 bytes signed integer max current> | Encoded value = max current * 100 | Encoded value = max current * 100 | LH | Example Set Max Current to 292mA: CwMA<29200><CRC> Hex: 0x43 0x77 0x4D 0x41 0x72 0x10 0x98 0xEC | <Prefix><Coding><Channel><DataOut><crc16><end> | "C" | "M" | "A" | <2 bytes signed integer max current> | Max current = encoded value / 100 | Max current = encoded value / 100 | LH | "\r\n" | CMA<encoded max current><CRC> \r\n | This value is also used to encode the current within the firmware.
03-Lens Commands and Properties | Get Max Output Current | No | ACTIVE | A,B,C,D,E | 1.8 | #0401 | Get the maximum firmware output current in channel A. | yes | <Prefix><Read><Coding><Channel><DataIn><crc16> | "C" | "r" | "M" | "A" | <2 dummy bytes> |  |  | LH | Example Get the Max Current: CrMA<0><0><CRC> Hex: 0x43 0x72 0x4D 0x41 0x00 0x00 0x71 0x80 | <Prefix><Coding><Channel><DataOut><crc16><end> | "C" | "M" | "A" | <2 bytes signed integer max current> | Max current = encoded value / 100 | Max current = encoded value / 100 | LH | "\r\n" | CMA<encoded max current><CRC> \r\n
03-Lens Commands and Properties | Set Focal Power | No | ACTIVE | A,B,C,E,F | 1.8 | #0310 | Set the focal power in channel A. Command can only be run under Controlled mode (command #0308). Focal power has an effective range of -5 to 15.48 diopters. | No | <Prefix><Write><Coding><Channel><DataIn><crc16> | "P" | "w" | "D" | "A" | <2 bytes unsigned integer focal power><2 dummy bytes> | Encoded value = (focal power + 5) * 200 | Encoded value = (focal power ) * 200 | LH | Example Set Focal Power to 1 diopter: PwDA<1200><0><0><CRC> Hex: 0x50 0x77 0x44 0x41 0x04 0xB0 0x00 0x00 0x31 0xA7 | none | none | none | none | none | none | none | none | none | none | If the focal power is set in control mode the focal power will be maintained on the lens by changing the current depending on the temperature. In any other mode, if the focal power is set by the user using this command, it will be changed interally however this value will not be used to maintain any property on the lens.
03-Lens Commands and Properties | Get Focal Power | No | ACTIVE | A,B,C,E,F | 1.8 | #0310 | Get the focal power in channel A. Command can only be run under Focal Power Controlled mode (command #0308). Focal power range is from -5 to 15.48 diopters. | yes | <Prefix><Read><Coding><Channel><DataIn><crc16> | "P" | "r" | "D" | "A" | <4 dummy bytes> |  |  | LH | Example Get Focal Power: PrDA<0><0><0><0><CRC> Hex: 0x50 0x72 0x44 0x41 0x00 0x00 0x00 0x00 0x64 0x70 | <Prefix><Coding><DataOut><crc16><end> | "P" | "D" | none | <2 bytes signed integer focal power> | Focal power = (encoded value/200) -5 | Focal power = (encoded value/200) -5 | LH | "\r\n" | PD<encoded focal power><CRC>\r\n
03-Lens Commands and Properties | Write byte to EEPROM | No | NEW | A,B,C,D,E | 1.9 | #9998 | Use to write the Temp Sensor using the Lens Driver. | yes | <Prefix><DataIn><crc16> | "Z" | "w" | none | none | <1 byte unsigned integer Register ID><1 byte unsigned integer byte to write> |  |  | LH | Example Write 1 to EEPROM byte 10: Zw<10><1><CRC> Hex: 0x5A 0x77 0x0A 0x01 0x65 0xA2 | <Prefix><DataOut><crc16><end> | "Z" | none | none | <1 byte unsigned integer Written Byte > |  |  | LH | "\r\n" | Z<written byte><CRC>\r\n | This command writes a new set of data to the lens EEPROM. When using this command ensure that the new set of data being loaded to the lens EEPROM is valid. If the set of data written to the lens EEPROM is invalid then the firmware will not work in control mode and the lens EEPROM will need to re-written.
03-Lens Commands and Properties | Read address from the EEPROM | No | NEW | A,B,C,D,E | 1.9 | #0609 | Read a address from the EEPROM | yes | <Prefix><coding><DataIn><crc16> | "Z" | "r" | none | none | <1 byte unsigned integer> |  |  | LH | Example read address EEPROM: Zr<1><CRC> Hex: 0x5A 0x72 0x0A 0x84 0xB4 | <Prefix><DataOut><crc16><end> | "Z" | none | none | <1 byte unsigned integer> |  |  | LH | "\r\n" | Z<100><CRC>\r\n
03-Lens Commands and Properties | Read EEPROM | No | ACTIVE | A,B,C,D,E | 1.8 | #9999 | Dump all the EEPROM. | yes | <Prefix><DataIn><crc16> | "D" | none | none | none | <1 dummy byte> |  |  | LH | Read EEPROM read Command: D<0><0><CRC> Hex: 0x44 0x00 0x00 0x40 0x15 | <Prefix><DataOut><end> | "D" | none | none | 256 Bytes from the temp sensor EEPROM |  |  | LH | "\r\n" | D<temp sensor eeprom>\r\n | This command reads the lens EEPROM and returns the entire set of data on the lens EEPROM.
04-Operation Mode | Change to Sinusoidal Signal | No | ACTIVE | A,B,C,D,E | 1.8 | #0301 | Sinusoidal Signal - Set channel A to Sinusoidal Waveform | yes | <Prefix><Write><Coding><Channel><crc16> | "M" | "w" | "S" | "A" | none |  |  | LH | Example Sinusoidal Mode Command: MwSA<CRC> Hex: 0x4D 0x77 0x53 0x41 0x5B 0xB6 | <Prefix><Coding><Channel><crc16><end> | "M" | "S" | "A" | none |  |  | LH | "\r\n" | MSA<CRC> \r\n
04-Operation Mode | Sinusoidal Signal (read state) | No | DEPRECATED | A,B,C,D,E | 1.8 | #0301 | Sinusoidal Signal - Set channel A to Sinusoidal Waveform DEPRECATED. This command will be remove in the next release. | yes | <Prefix><Read><Coding><Channel><crc16> | "M" | "r" | "S" | "A" | none |  |  | LH | Example Sinusoidal Mode Command: MrSA<CRC> Hex: 0x4D 0x72 0x53 0x41 0x4B 0xB7 | <Prefix><Coding><Channel><crc16><end> | "M" | "S" | "A" | none |  |  | LH | "\r\n" | MSA<CRC> \r\n | The command will return the actual mode, not necessarily MSA<CRC>
04-Operation Mode | Change to Rectangular Signal | No | ACTIVE | A,B,C,D,E | 1.8 | #0302 | Square Signal - Set channel A to square Waveform | yes | <Prefix><Write><Coding><Channel><crc16> | "M" | "w" | "Q" | "A" | none |  |  | LH | Example Square Mode Command: MwQA<CRC> Hex: 0x4D 0x77 0x51 0x41 0x5A 0xD6 | <Prefix><Coding><Channel><crc16><end> | "M" | "Q" | "A" | none |  |  | LH | "\r\n" | MQA<CRC> \r\n
04-Operation Mode | Rectangular Signal (read state) | No | DEPRECATED | A,B,C,D,E | 1.8 | #0302 | Square Signal - Set channel A to square Waveform DEPRECATED. This command will be remove in the next release. | yes | <Prefix><Read><Coding><Channel><crc16> | "M" | "r" | "Q" | "A" | none |  |  | LH | Example Square Mode Command: MrQA<CRC> Hex: 0x4D 0x72 0x51 0x41 0x4A 0xD7 | <Prefix><Coding><Channel><crc16><end> | "M" | "Q" | "A" | none |  |  | LH | "\r\n" | MQA<CRC> \r\n | The command will return the actual mode, not necessarily MQA<CRC>
04-Operation Mode | Change to Current mode | No | ACTIVE | A,B,C,D,E | 1.8 | #0303 | Current Mode - Set channel A to current mode | yes | <Prefix><Write><Coding><Channel><crc16> | "M" | "w" | "D" | "A" | none |  |  | LH | Example Current Mode Command: MwDA<CRC> Hex: 0x4D 0x77 0x44 0x41 0x54 0x46 | <Prefix><Coding><Channel><crc16><end> | "M" | "D" | "A" | none |  |  | LH | "\r\n" | MDA<CRC> \r\n
04-Operation Mode | Current mode (read state) | No | DEPRECATED | A,B,C,D,E | 1.8 | #0303 | Current Mode - Set channel A to current mode DEPRECATED. This command will be remove in the next release. | yes | <Prefix><Read><Coding><Channel><crc16> | "M" | "r" | "D" | "A" | none |  |  | LH | Example Current Mode Command: MrDA<CRC> Hex: 0x4D 0x72 0x44 0x41 0x44 0x47 | <Prefix><Coding><Channel><crc16><end> | "M" | "D" | "A" | none |  |  | LH | "\r\n" | MDA<CRC> \r\n | The command will return the actual mode, not necessarily MDA<CRC>
04-Operation Mode | Change to Triangular Signal | No | ACTIVE | A,B,C,D,E | 1.8 | #0304 | Triangular Signal - Set channel A to square Waveform | yes | <Prefix><Write><Coding><Channel><crc16> | "M" | "w" | "T" | "A" | none |  |  | LH | Example Triangular Command: MwTA<CRC> Hex: 0x4D 0x77 0x54 0x41 0x59 0x86 | <Prefix><Coding><Channel><crc16><end> | "M" | "T" | "A" | none |  |  | LH | "\r\n" | MTA<CRC> \r\n
04-Operation Mode | Triangular Signal (read state) | No | DEPRECATED | A,B,C,D,E | 1.8 | #0304 | Triangular Signal - Set channel A to square Waveform DEPRECATED. This command will be remove in the next release. | yes | <Prefix><Read><Coding><Channel><crc16> | "M" | "r" | "T" | "A" | none |  |  | LH | Example Triangular Command: MrTA<CRC> Hex: 0x4D 0x72 0x54 0x41 0x49 0x87 | <Prefix><Coding><Channel><crc16><end> | "M" | "T" | "A" | none |  |  | LH | "\r\n" | MTA<CRC> \r\n | The command will return the actual mode, not necessarily MTA<CRC>
04-Operation Mode | Change to Analog Mode | No | ACTIVE | A,B,C,E | 1.8 | #0321 | Enables Analog Mode. Analog Mode controls the current output based on the analog reading from the anlog input pin. | yes | <Prefix><Write><Coding><Channel><crc16> | "M" | "w" | "A" | "A" | none |  |  | LH | Example Analog Command: MwAA<CRC> Hex: 0x4D 0x77 0x41 0x41 0x57 0x16 | <Prefix><Coding><Channel><crc16><end> | "M" | "A" | "A" | none |  |  | LH | "\r\n" | MAA<CRC> \r\n | If the current is set by the user in Analog Mode it will be overwritten by the reading from the anlog input pin. If the user wishes to change the current output, the analog input value must be changed or the mode must be changed.
04-Operation Mode | Analog Mode (read state) | No | DEPRECATED | A,B,C,E | 1.8 | #0321 | Enables Analog Mode. Analog Mode controls the current output based on the analog reading from the anlog input pin. DEPRECATED. This command will be remove in the next release. | yes | <Prefix><Read><Coding><Channel><crc16> | "M" | "r" | "A" | "A" | none |  |  | LH | Example Analog Command: MrAA<CRC> Hex: 0x4D 0x72 0x41 0x41 0x47 0x17 | <Prefix><Coding><Channel><crc16><end> | "M" | "A" | "A" | none |  |  | LH | "\r\n" | MAA<CRC> \r\n | If the current is set by the user in Analog Mode it will be overwritten by the reading from the anlog input pin. If the user wishes to change the current output, the analog input value must be changed or the mode must be changed. The command will return the actual mode, not necessarily MAA<CRC>
04-Operation Mode | Change to Focal Power Controlled Mode | No | ACTIVE | A,B,C,D,E | 1.8 | #0308 | Enter focal power controlled Mode - Allow channel A to maintain a constant lens focal power. This command requires temperature limits to be set (Command #0309). In addition, the command responds with the achievable focal power range at the given temperature limits, and the latest focal power. | yes | <Prefix><Write><Coding><Channel><crc16> | "M" | "w" | "C" | "A" | none |  |  | LH | Example Focal Power Controlled Mode Command: MwCA<CRC> Hex: 0x4D 0x77 0x43 0x41 0x56 0x76 | <Prefix><Coding><Channel><DataOut><crc16><end> | "M" | "C" | "A" | <1 bytes unsigned integer ErrorByte><2 bytes unsigned integer max focal power controlled limit><2 bytes unsigned integer min focal power controlled limit> | Focal power = (encoded value/200) -5 | Focal power = encoded value/200 | LH | "\r\n" | MCA<errorByte><encoded max focal power limit><encoded min focal power limit><CRC> \r\n | Error byte defniition see sheet "Help". This mode can only be entered if there is a valid EEPROM for interpolation calculation.
04-Operation Mode | Focal power controlled mode (read state) | No | DEPRECATED | A,B,C,D,E | 1.8 | #0308 | Enter focal power controlled Mode - Allow channel A to maintain a constant lens focal power. This command requires temperature limits to be set (Command #0309). In addition, the command responds with the achievable focal power range at the given temperature limits, and the latest focal power. DEPRECATED. This command will be remove in the next release. | yes | <Prefix><Read><Coding><Channel><crc16> | "M" | "r" | "C" | "A" | none |  |  | LH | Example Focal Power Controlled Mode Command: MrCA<CRC> Hex: 0x4D 0x72 0x43 0x41 0x46 0x77 | <Prefix><Coding><Channel><DataOut><crc16><end> | "M" | "C" | "A" | <1 bytes unsigned integer ErrorByte><2 bytes unsigned integer max focal power controlled limit><2 bytes unsigned integer min focal power controlled limit> | Focal power = (encoded value/200) -5 | Focal power = encoded value/200 | LH | "\r\n" | MCA<errorByte><encoded max focal power limit><encoded min focal power limit><CRC> \r\n | Error byte defniition see sheet "Help". This mode can only be entered if there is a valid EEPROM for interpolation calculation. The command will return the actual mode, not necessarily MCA<CRC>
04-Operation Mode | Read Curently Active Mode | No | NEW | A,B,C,D,E | 1.9 | #0312 | This command will return the actual mode. | yes | <Prefix><Coding><Channel><crc16> | "M" | none | "M" | "A" | none |  |  | LH | Example Read Mode Command: MMA<CRC> Hex: 0x4D 0x4D 0x41 0x58 0x6B | <Prefix><Coding><Channel><DataOut><crc16><end> | "M" | "M" | "A" | <1byte unsigned integer> - Current Mode = 1 - Sinusoidal Signal Mode = 2 - Triangular Mode = 3 - Retangular Mode = 4 - Focal Power Mode = 5 - Analog Mode = 6 - Position Controlled Mode = 7 |  |  | LH | "\r\n" | MMA<mode><CRC> \r\n
06-Advanced Commands | Set Lens Driver to DFU Mode | No | ACTIVE | A,B,C,D,E | 1.8 | #0801 | Restart the device in DFU mode. | no | <Prefix><crc16> | "RESET" | none | none | none | none |  |  | LH | Example Rest Command: RESET<CRC> Hex: 0x52 0x45 0x53 0x45 0x54 0x6F 0xBE | none | "RESET" | none | none | none |  |  | LH | "\r\n" |  | This command is used to reset the firmware. After this command is used another version of the firmware must be loaded to the lens dirver before the lens driver will function normally
06-Advanced Commands | Interpoaltion Command | No | ACTIVE | A,B,C,D,E | 1.8 | #9997 | Returns the equivalent focal power for a given temperature and current using the interpolation method that is used for control mode | yes | <Prefix><DataIn><crc16> | "Y" | none | none | none | <2 bytes signed integer current><2 bytes signed integer temperature> | Encoded value (current) = current * (4095 /max current), where the default max current is 293mA (command #0401). Encoded current values can only be within the range -4095 to 4095 Encoded value (temperature) = temperature * 256 | Encoded value (current) = current * (4095 /max current), where the default max current is 293mA (command #0401). Encoded current values can only be within the range -4095 to 4095 Encoded value (temperature) = temperature * 256 | LH | Example Interpolation Command for currnet of 100mA and Temperature 20'C: Y<1402><5120><CRC> Hex: 0x59 0x07 0x5A 0x14 0x00 0x32 0x6A | <Prefix><DataOut><crc16><end> | "Y" | none | none | <2 bytes signed integer equivalent focal power> | Focal power = (encoded value/200) -5 | Focal power = encoded value/200 | LH | "\r\n" | Y<16 bit signed integer equivalent focal power>\r\n | This command can be used to find what the firmware will return as the focal power for a given current and temperature. This command can be used for testing purposes.

## CRC
CRC-16-IBM Calculation
Written in C#:
Polynomial: x^16 + x^15 + x^2 + 1 (0xa001)
Initial value: 0x0000
public class CRC16IBM
{
const ushort polynomial = 0xA001;
ushort[] table = new ushort[256];
public ushort ComputeChecksum(byte[] bytes)
{
ushort crc = 0;
for (int i = 0; i < bytes.Length; ++i)
{
byte index = (byte)(crc ^ bytes[i]);
crc = (ushort)((crc >> 8) ^ table[index]);
}
return crc;
}
public byte[] ComputeChecksumBytes(byte[] bytes) {
ushort crc = ComputeChecksum(bytes);
return BitConverter.GetBytes(crc);
}
public CRC16IBM()
{
ushort value;
ushort temp;
for(ushort i = 0; i < table.Length; ++i) {
value = 0;
temp = i;
for(byte j = 0; j < 8; ++j) {
if(((value ^ temp) & 0x0001) != 0) {
value = (ushort)((value >> 1) ^ polynomial);
}else {
value >>= 1;
}
temp >>= 1;
}
table[i] = value;
}
}
}

## Example for 0 & 3 diopters
Example
How to connect to Lens Driver 4 and set the focal power to 0 and 3 diopters
1) | Initialize Driver
| String | Start
2) | Set Temperature Limits
| Value | P | w | T | A | 20°C |  | 35°C |  | CRC |  | Hex String
[6 empty] | <upper Temperature high byte> | <upper temperature low byte> | <lower Temperature high byte> | <lower temperature low byte> | <CRC high byte> | <CRC low byte>
| Hex | 0x50 | 0x77 | 0x54 | 0x41 | 0x02 | 0x30 | 0x01 | 0x40 | 0x32 | 0x37 | 50775441023001403237
3) | Set Focal Power Mode
| Value | M | w | C | A | CRC | [5 empty] | Hex String
[6 empty] | <CRC high byte> | <CRC low byte>
| Hex | 0x4d | 0x77 | 0x43 | 0x41 | 0x56 | 0x76 |  |  |  |  | 4d7743415676
4) | Set Focal Power: 0dpt
| Value | P | w | D | A | 0 dpt = 1000 |  | dummy data |  | CRC |  | Hex String
[6 empty] | <focal power high byte> | <focal power low byte> | <dummy data high byte> | <dummy data low byte> | <CRC high byte> | <CRC low byte>
| Hex | 0x50 | 0x77 | 0x44 | 0x41 | 0x03 | 0xe8 | 0x00 | 0x00 | 0xb1 | 0x00 | 5077444103e80000b100
5) | Set Focal Power: 3dpt
| Value | P | w | D | A | 3 dpt = 1600 |  | dummy data |  | CRC |  | Hex String
[6 empty] | <focal power high byte> | <focal power low byte> | <dummy data high byte> | <dummy data low byte> | <CRC high byte> | <CRC low byte>
| Hex | 0x50 | 0x77 | 0x44 | 0x41 | 0x06 | 0x40 | 0x00 | 0x00 | 0x30 | 0x2c | 5077444106400000302c

## Example for read commands
Example
Typical read-out commands
1) | Initialize Driver
| String | Start
2) | Send GetTemperature command
| Value | T | A | CRC | [6 empty] | Hex String
|  |  |  | <CRC high byte> | <CRC low byte>
| Hex | 0x54 | 0x41 | 0xFE | 0xF0 | [5 empty] | 5441FEF0
3) | Decode reply
| Value | T | A | Error byte | <16 bit signed integer temperature> |  | CRC |  | \r | \n | Hex String
[7 empty] | <CRC high byte> | <CRC low byte>
| Hex | 0x54 | 0x41 | 0x00 | 0x01 | 0xB2 | 0xA4 | 0xFF | 0x0D | 0x0A | 54410001B2A4FF0D0A
[5 empty] | Integer: | 434
[5 empty] | Temperature: | 27.125 | °C
4) | Set current to 100mA
| Value | A | w | <16 bit signed integer current |  | CRC |  |  |  |  | Hex String
[6 empty] | <CRC high byte> | <CRC low byte>
| Hex | 0x41 | 0x77 | 0x05 | 0x7A | 0x26 | 0x95 |  |  |  | 4177057A2695
5) | Send GetCurrent command
| Value | A | r | Dummy bytes |  | CRC |  |  |  |  | Hex String
|  |  |  | <CRC high byte> | <CRC low byte> | <CRC high byte> | <CRC low byte>
| Hex | 0x41 | 0x72 | 0x00 | 0x00 | 0xb4 | 0x27 |  |  |  | 41720000b427
6) | Decode reply
| Value | A | <16 bit signed integer current |  | CRC |  | \r | \n |  |  | Hex String
[5 empty] | <CRC high byte> | <CRC low byte>
| Hex | 0x41 | 0x05 | 0x7A | 0xD2 | 0xFF | 0x0D | 0x0A |  |  | 41057AD2FF0D0A
|  |  | Integer: | 1402
|  |  | Current | 100.2896 | mA

## Video
Video
Short tutorial on simple handshake with Lens Driver 4 using Hterm hyperterminal
http://youtu.be/ms4RCm5hcUY